## 1. Configuration
CARBELIM PRIVATE LIMITED
AirForest™ Lab-Grade Photobioreactor — Product Configuration & Quotation Sheet
Version 1.0 ¦ Coimbatore, India ¦ IIT Madras Research Park Incubatee
Client / Organisation |  | Dept. of Marine Biotechnology, Naval Materials Research Laboratory
Contact Person |  | Enter name
Email / Phone |  | Enter contact details
Quotation Date |  | DD-MMM-YYYY
Quotation Ref # |  | CBL-PBR-001
Delivery Location |  | Dept. of Marine Biotechnology, Naval Materials Research Laboratory, Ambernath (Maharashtra)
CONFIGURATION OPTIONS — Tick / Highlight your selection in yellow cells
Module | Option A | Option B | Option C | Notes
Air Sensor Coverage | Inlet Only | Inlet + Outlet
Connectivity | WiFi Only (customer's router) | WiFi + 4G SIM (standalone)
IoT Dashboard | None | Cloud Dashboard (AWS-hosted)
Electrical Supply | 240V 50/60Hz | 230 V / 50 Hz (Indian/International)
PRICE SUMMARY (see Sheet 2 for full BOM)
STRUCTURAL — Frame & Tank |  |  |  | 45750
LIGHTING — LED Grow Lights (Side-Mounted) |  |  |  | 6500
LIQUID SENSORS — pH, Turbidity, TDS, Temperature, Level |  |  |  | 21500
AIR SENSORS — Option A |  |  |  | 16200
ELECTRICAL — RCCB, MCBs, Relays, Wiring |  |  |  | 10500
CONTROL SYSTEM — PLC + HMI (IDEC FT2J-7U) |  |  |  | 35000
MISCELLANEOUS & CONSUMABLES |  |  |  | 8000
Connectivity Module (as configured) |  |  |  | #REF!
IoT Dashboard (annual, if selected) |  |  |  | 29000
Subtotal (INR) |  |  |  | #REF!
GST @ 18% |  |  |  | #REF!
GRAND TOTAL (INR) |  |  |  | #REF!
Note: Prices are indicative (INR, ex-works Coimbatore). Taxes, freight, installation, and commissioning charged separately. Validity: 30 days from quotation date.

## 2. BOM & Pricing
CARBELIM PRIVATE LIMITED — AirForest™ PBR ¦ Full Bill of Materials & Pricing
Item # | Category | Component / Description | Part Number / Model | Qty | Unit | Unit Price (INR) | Per Unit Margin | Total (INR) | Total Margin | Remarks
STRUCTURAL — Frame & Tank
1 | STRUCTURAL | SS304 Matte Frame | Custom Fabricated | 1 | Set | 12750 | 26775 | 12750 | 26775 | Includes base feet & mounting brackets
2 | STRUCTURAL | Acrylic Tank (50 L Capacity) | Custom Fabricated | 1 | No. | 30000 | 63000 | 30000 | 63000 | Silicone-sealed joints; drain port
3 | STRUCTURAL | Stainless Steel Hardware (bolts, nuts, standoffs) | SS-304 Assorted | 1 | Lot | 3000 | 6300 | 3000 | 6300
Subtotal — Frame & Tank | [7 empty] | 45750 | 96075
LIGHTING — LED Grow Lights (Side-Mounted)
5 | LIGHTING | Full-Spectrum LED Grow Strip (side panels, 4 sides) | CARB LED equiv. | 5 | Strip | 1000 | 2800 | 5000 | 14000
6 | LIGHTING | LED Driver / Power Supply 24 V (for lighting) | Power Supply | 1 | No. | 1500 | 4200 | 1500 | 4200 | 100 W, 24 V / 4.2 A, DIN-rail
Subtotal — LED Grow Lights (Side-Mounted) | [7 empty] | 6500 | 14950
LIQUID SENSORS — pH, Turbidity, TDS, Temperature, Level
8 | WATER SENSORS | pH Sensor + Transmitter | Carbelim CBPH01 | 1 | No. | 8000 | 22400 | 8000 | 22400 | 0–14 pH, 0–60 °C; BNC probe; 0–5 V output
9 | WATER SENSORS | Turbidity Sensor (Basic) | Carbelim CBTURB01 | 1 | No. | 1600 | 4480 | 1600 | 4480 | 0–3000 NTU; 5 V DC; analog 0–4.5 V output
10 | WATER SENSORS | TDS Sensor (Basic) | Carbelim CBTDS01 | 1 | No. | 1400 | 3920 | 1400 | 3920 | 0–1000 ppm; 3.3–5.5 V; analog 0–2.3 V output
11 | WATER SENSORS | Water Temperature — PT100 RTD + Transmitter | Carbelim CBPT100 | 1 | No. | 1300 | 3640 | 1300 | 3640 | -10 to 100 °C; SPI interface to PLC/MCU
12 | WATER SENSORS | Liquid Level — Pressure Sensor | Carbelim LEVEL01 | 1 | No. | 2000 | 5600 | 2000 | 5600 | 4–20 mA output; submersible; stainless diaphragm
13 | WATER SENSORS | Signal Converter 0–5 V → 4–20 mA (DIN rail, ×3) | Carbelim Signal multiplexer | 1 | No. | 1200 | 3360 | 1200 | 3360 | For pH, turbidity, TDS long cable runs to PLC/PCB
14 | Sensor Enclosure | Enclosure for Wate Sensors | Carbelim SENENCLO | 1 | No. | 6000 | 12300 | 6000 | 12300 | Enclsoure for Electrodes porbes
Subtotal — pH, Turbidity, TDS, Temperature, Level | [7 empty] | 21500 | 55700
AIR SENSORS — Option A: Inlet Only
15 | AIR SENSORS | Multi-Gas Air Quality Module — INLET | Carbelim CBZPHS01B | 1 | No. | 15000 | 48000 | 15000 | 48000 | CO₂, CO, O₃, TVOC, HCHO, Temp, RH, PM2.5, PM10; RS-485 Modbus
16 | AIR SENSORS | Sensor Enclosure + Mounting Bracket (Inlet) | Custom RS485 Hub | 1 | No. | 1200 | 2760 | 1200 | 2760 | RS-485 cable entry
Subtotal — Option A: Inlet Only | [7 empty] | 16200 | 50760
AIR SENSORS — Option B: Inlet + Outlet
17 | AIR SENSORS | Multi-Gas Air Quality Module — OUTLET | Carbelim CBZPHS01B | 2 | No. | 15000 | 48000 | 30000 | 96000 | CO₂, CO, O₃, TVOC, HCHO, Temp, RH, PM2.5, PM10; RS-485 Modbus
18 | AIR SENSORS | Sensor Enclosure + Mounting Bracket (Outlet) | Custom RS485 Hub | 1 | No. | 1200 | 2760 | 1200 | 2760 | RS-485 cable entry
Subtotal — Option B: Inlet + Outlet (additional to D-A) | [7 empty] | 31200 | 98760
ELECTRICAL — RCCB, MCBs, Relays, Wiring, Terminals
19 | Protection | Circuit Breakers (Main RCCB & Branch MCBs) | Main RCCB & Branch MCBs | 1 | Lot | 2800 | 12600 | 2800 | 12600 | Main earth-leakage and circuit overload protection
20 | Air Compression unit | Carbelim CAQ-013 | 4 Watts, 2*4LPM Aeration | 1 | No. | 600 | 2700 | 600 | 2700 | Aeration
21 | Intake Auxiliary fan | Carbelim CIAF | 12V DC/ 240V AC | 1 | No. | 600 | 2700 | 600 | 2700
22 | Enclosure, Connection and Control | Waterproof Switchboard, Wiring Kit (Cables, Terminals, DIN Rails, Ducts), Relays (Solid-State & DIN-Rail), Power Supplies (24V/12V SMPS) & Power Inlet | Waterproof Switchboard, Cables, Terminals, DIN Rails, Ducts, SSR and Relays, Cables, Terminals, DIN Rails, Ducts, SMPS (24V/12V SMPS) | 1 | No. | 6500 | 29250 | 6500 | 29250 | Safe housing for all electrical components, Power routing, sensor wiring, and component mounting
Subtotal — RCCB, MCBs, Relays, Wiring, Terminals | [7 empty] | 10500 | 47250
CONTROL SYSTEM — PLC + HMI (IDEC FT2J-7U)
23 | CONTROL SYSTEM | HMI Logic Control Board | Carbelim IDEC PLC (FT2J-7U22RAF-B) | 1 | No. | 35000 | 161000 | 35000 | 161000 | Analog DIgital Input, 24 V DC
Subtotal — HMI + PLC (IDEC FT2J-7U) | [7 empty] | 35000 | 161000
CONNECTIVITY — Option B: WiFi + 4G SIM (additional)
24 | CONNECTIVITY | 4G LTE + WiFi Module | SIM Card based Router for WiFi/ Ethernet | 1 | No. | 1500 | 3450 | 1500 | 3450 | LTE Cat-4; dual SIM;
Subtotal — Option B: WiFi + 4G SIM (additional) | [7 empty] | 1500 | 3450
IoT DASHBOARD — Optional Cloud Service
25 | IoT DASHBOARD | AWS IoT Core + DynamoDB + S3 Setup (one-time) | Carbelim Cloud Setup | 1 | Service | 25000 | 100000 | 25000 | 100000 | One-time config; real-time MQTT + REST API
27 | IoT DASHBOARD | Historical Analytics Module (per device/year) | Carbelim SaaS CBL-HIST-01 | 1 | Year | 4000 | 10000 | 4000 | 10000 | Athena-backed long-term trend analysis + reports
Subtotal — Optional Cloud Service | [7 empty] | 29000 | 110000
MISCELLANEOUS & CONSUMABLES
28 | MISCELLANEOUS & CONSUMABLES | Cable Labels + Ferrule Kit, Zip Ties, Heat Shrink, Electrical Tape (lot), Documentation Pack (wiring diagrams, BOM, manual) | Partex / Brady assorted | 1 | Lot | 8000 | 18400 | 8000 | 18400
Subtotal — MISCELLANEOUS & CONSUMABLES | [7 empty] | 8000 | 18400

## 3. Technical Specs
CARBELIM AirForest™ — Lab-Grade PBR ¦ Technical Specifications
Parameter | Specification / Value | Notes
1. MECHANICAL & STRUCTURAL
Overall Dimensions (W × D × H) | TBD | Varies with tank volume
Frame Material | MS Square Tube
Tank Material | Cast Acrylic
Tank Volume (nominal) | 50 litres
Operating Weight (filled) | TBD | Incl. frame, tank, electronics
Ingress Protection — Panel | Electrical enclosure
2. LIGHTING
Light Type | Full-Spectrum LED, side-mounted
Number of Panels | 4 sides (full perimeter)
Power per Strip | TBD | TBD
Control | Relay based Control | 0–10 V analog output
Voltage | 24 V DC | Powered by Power Supply
3. LIQUID SENSING
pH Range | 0–14 pH | Carbelim CBPH01; ±0.1 pH accuracy
Turbidity Range (Basic) | 0–3000 NTU | Carbelim CBTURB01; ±5% FS
TDS Range (Basic) | 0–1000 ppm | Carbelim CBTDS01; ±10% FS
Water Temperature | -10 to 100 °C | Carbelim CBPT100; ±0.5 °C
Liquid Level | 0–500 mm (hydrostatic) | Carbelim LEVEL01 4-20ma; ±1% FS
DO Sensor | 0 to 20.0 ppm (mg/L) | Carbelim DOSensor ±0.2 ppm to ±1 count
4. AIR SENSING (Carbelim CBZPHS01B)
Parameters Measured | CO₂, CO, O₃, TVOC, HCHO, Temp, RH, PM2.5, PM10 | 11 Modbus registers
CO₂ Range | 400–5000 ppm | ±50 ppm or ±5% of reading
PM2.5 / PM10 Range | 0–999 µg/m³ | Laser scattering method
Temperature (air) | -10 to 50 °C | ±0.5 °C
Relative Humidity | 0–100% RH | ±3% RH
Communication | RS-485 Modbus RTU, 9600 baud | FC03; addresses 0–10
Sensor Positions | Inlet (Option A) or Inlet + Outlet (Option B) | Option B adds 1 ZPHS01B
5. CONTROL SYSTEM
PCB / HMI + PLC | IDEC FT2J-7U/ Carbelim PCB | 7" TFT touch; RS-485; Ethernet
6. CONNECTIVITY & IoT
WiFi Option | IEEE 802.11 b/g/n; 2.4 GHz | Connect to existing LAN router
SIM Option (add-on) | 4G LTE Cat-4 | Standalone; nano SIM
Protocol | MQTT over TLS (AWS IoT Core) | QoS 1; retained messages
Cloud Backend | AWS IoT Core + DynamoDB + S3 | Serverless; scales to fleet
Dashboard | Web-based (responsive); WebSocket real-time | Mobile + desktop; no app install
Data Retention | Hot: 30 days DynamoDB; Cold: 2 years S3 | Tiered TTL strategy
7. ELECTRICAL
Supply Voltage | 230 V AC / 50 Hz (India) | Specify at order
Total Connected Load | ~300 W typical (excl. heater) | ~550 W with 250 W EHEIM heater
Main Protection | RCCB 2P 30 mA + MCB 2P 16 A | Circuit breaker
Heater Control | SSR 40 A | Circuit breaker
SMPS Outputs | 24 V DC (sensors + relays) + 12 V DC (PCB) | 24v 150w Power Supply
Air Compression unit | 4 Watts, 2*4LPM Aeration
Specifications subject to change without notice. All prices are estimates only. Contact Carbelim for final engineering quotation.